## Slide 1: Modern imaging techniques in biology
The physical background of medical tomographies
Lecture 1
Image: www.123rf.com

## Slide 2: How to get a mark?
Alternatives:
Oral exam
Mini review based on 3 scientific papers: ~10 pages. Consultation needed.

## Slide 3: Literature
You can find English literature at my website:
http://balintszabo.web.elte.hu/modernkepalkototechnikak.html

## Slide 4: Thematics of the 10 lectures in the semester
X-ray CT (computed tomography) scanner: 2 lectures.
MRI and fMRI: (functional) magnetic resonance imaging: 6 lectures.
PET: positron emission tomography: 1 lecture.
Sonography: medical ultrasound: 1 lecture.
Visit to a CT-MRI diagnostic center in Budapest in the end of the semester.
Image: BBC news, 2001 January.

## Slide 5: Detailed thematics: X-ray CT scanner
Principles of CT measurement. Attenuation scale of human tissues. Parallel beam, ring-detector and fan-beam systems. Beam hardening.
Imaging. Mathematical background: Radon transformation, Point Spread Function (PSF), convolution, filter equation in 2D and 1D.
Image: http://www.bwequinevets.co.uk

## Slide 6: Detailed thematics: MRI
Microscopic and macroscopic magnetization. The Bloch equation. T1 and T2 relaxation times.
Magnetic resonance. 90° pulse. FID: free induction decay. Effective T2.
Pulse method for determining T2 and T1. Spin echo. Magnetic resonance spectroscopy.
Imaging. Selective excitation and read out of a slice.
Pulse sequences and contrast. Flow in MRI.
fMRI. BOLD. Echo planar imaging (EPI). Spin echo (SE) and gradient echo (GE) for fast imaging.
Image: http://radiologia.blog.hu

## Slide 7: X-ray CT scanner: history
The method produces axial transverse tomograms: body layers perpendicular to the longitudinal axis of the body.
History of CT
1963-64. A. M. Cormack: first to describe an X-ray tomography corresponding to CT.
1917. Mathematical solution of the problem had been given by Radon.
1972. First clinical application.
1979: Nobel prize for Sir Hounsfield and A. M. Comack
For now it has become a routine diagnostic method.
# of CT scans in the U.S. in a year:
  1980: 3 million
  1995: 20 million
  2006: 60 million
The prototype of CT scanner, Wikipedia.

## Slide 8: X-ray 2D shadow image
Classical X-ray: 2D shadow image projected by the radiation cone into the image plane. Image is created by superpositioning several object elements
Image:
https://www.slideshare.net/ihmei/reading-chest-xray

## Slide 9: X-ray CT scanner advantages vs classical shadow image
Image free of superpositioning: substitution image. This is the case of CT. Results in soft tissue images with high contrast.
High contrast can be obtained without using a contrast medium, e.g., white and gray cerebral matter can be distinguished easily. White matter appears darker due to its high fat content, i.e. lower attenuation capacity.
CT scan of a brain
A series of CT scans converted into an animated, Wikipedia

## Slide 10: Parallel beam, fan-beam, and ring-detector setups
Images: https://radiologykey.com/

## Slide 11: Backprojection and filtered backprojection
Source:http://www.dspguide.com
http://www.impactscan.org/

## Slide 12: Hounsfield units of human tissues
Source: http://www.people.vcu.edu
Source: https://radiologykey.com/
X-ray linear attenuation coefficient: μ.
Lamberts’ law of absorption: intensity (I) decreases exponentially with thickness (x):
I=I0e-μx

## Slide 13: X-ray beam hardening
The beam is not monochromatic: contains different energy components. Beam intensity in a homogeneous material as a function of path length (x):
(E)
E: energy, : linear attenuation coefficient.
In an inhomogeneous material:
(E)

## Slide 14: X-ray beam hardening
Softer (lower energy) X-ray is attenuated (absorbed) more strongly in tissues
Soft components are filtered out -> beam gets harder when passing through the tissue.
https://radprotection.wikispaces.com
Correction for beam hardening. CT book, p. 105.